## Sheet1
|  | n=Number, l=letter |  |  |  | Univ_modes: 1=2-pos_stops, 2=2-pos_noStop, 3=multipos |  |  | Univ_modes: 1=2-pos_stops, 2=2-pos_noStop, 3=multipos
2-pos EM2C (older) |  | 2-pos E2CA (newer) |  | Multi-pos_EMH_EMT |  | Universal_Older_<=CT |  |  | Universal_EQ+
/?<enter> | Displays list of valid commands | ?<enter> | Displays a list of valid commands | /?<enter> | Displays a list of valid commands | /?<enter> | 1, 2, 3 | Displays a list of valid commands | /?<enter> | 1, 2, 3 | Displays a list of valid commands
[9 empty] | AL<enter> | 2, 3 | Moves the actuator drive shaft to the reference position prior to valve installation.
[6 empty] | AM<enter> | 1, 2, 3 | Displays the current actuator mode | AM<enter> | 1, 2, 3 | Displays the current actuator mode
[6 empty] | AMn<enter> | 1, 2, 3 | Sets the actuator mode to [1] two position with stops, [2] two position without stops, or [3] multiposition | AMn<enter> | 1, 2, 3 | Sets the actuator mode to [1] two position with stops, [2] two position without stops, or [3] multiposition
CC<enter> | Sends the actuator to Position B | CC<enter> | Sends the actuator to Position B | CCnn<enter> | Sends the actuator counterclockwise to position nn (from 1 to NP) | CC<enter> | 1, 2 | Sends the actuator to Position B | CC<enter> | 1, 2 | Sends the actuator from Position A to Position B
|  | CNT<enter> | Sets or displays the move count, from 0 to 65,000 | [5 empty] | CC<enter> | 3 | Decrements the actuator one position. (For example, from position 5 to position 4)
[6 empty] | CCnn<enter> | 3 | Sends the actuator counterclockwise to position nn (from 1 to NP) | CCnn<enter> | 3 | Sends the actuator in the “negative” or “down” direction to position nn (from 1 to NP)
[6 empty] | CNT<enter> | 1, 2, 3 | Displays the current value in the actuation counter | CNT<enter> | 1, 2, 3 | Displays the current value in the actuation counter
[6 empty] | CNTnnnnn<enter> | 1, 2, 3 | Sets the actuation counter from 0 to 65535. (For example, to reset the counter, type CNT0<enter>) | CNTnnnnn<enter> | 1, 2, 3 | Sets the actuation counter from 0 to 65535. (For example, to reset the counter, type CNT0<enter>)
CP<enter> | Displays the current actuator position | CP<enter> | Displays the current actuator position | CP<enter> | Displays the current position | CP<enter> | 1, 2, 3 | Displays the current position | CP<enter> | 1, 2, 3 | Displays the current position
CW<enter> | Sends the actuator to Position A | CW<enter> | Sends the actuator to Position A |  |  | CW<enter> | 1, 2 | Sends the actuator to Position A | CW<enter> | 1, 2 | Sends the actuator from Position B to Position A
[9 empty] | CW<enter> | 3 | Increments the actuator one position. (For example, from position 4 to position 5)
|  |  |  | CWnn<enter> | Sends the actuator clockwise to position nn (from 1 to NP) | CWnn<enter> | 3 | Sends the actuator clockwise to position nn (from 1 to NP) | CWnn<enter> | 3 | Sends the actuator in the “positive” or “up” direction to position nn (from 1 to NP)
[6 empty] | DELn<enter> | 1, 2, 3 | Sets the command delimiter to [1] CR (default), [2] LF, or [3] CRLF
DT<enter> | Displays the current delay time in milliseconds |  |  |  |  | DT<enter> | 1, 2 | Displays the current delay time in milliseconds | DT<enter> | 1, 2 | Displays the current delay time in milliseconds
DTn<enter> | Sets the delay time from 0 to 65,000 milliseconds NOTE: The total delay time equals n = 2 milliseconds | DT(n)<enter> | Sets or displays the delay time, where n = 0 to 65,535 milliseconds |  |  | DTnnnnn<enter> | 1, 2 | Sets the delay time from 0 to 65,000 milliseconds | DTnnnnn<enter> | 1, 2 | Sets the delay time from 0 to 65,000 milliseconds
GOn<enter> | Sends the actuator to Position n, where n is A or B | GO(l)<enter> | Sends the actuator to Position l, where l is A or B | GOnn<enter> | Sends the actuator to position nn (from 1 to NP) via the shortest route | GOnn<enter> | 1, 2 | Sends the actuator to position n, where n is A or B | GOnn<enter> | 1, 2 | Sends the actuator to position n, where n is A or B
[6 empty] | GOnn<enter> | 3 | Sends the actuator to position nn (from 1 to NP) via the shortest route | GOnn<enter> | 3 | Sends the actuator to position nn (from 1 to NP) via the shortest route
[6 empty] | HM<enter> | 3 | Moves the valve to postion 1 (home) | HM<enter> | 3 | Moves the valve to postion 1 (home)
|  | GV<enter> | Displays the measured DC input voltage
|  | GVD | Displays the measured DC input voltage drop while moving
ID*<enter> | Clears the ID variable | ID(l)<enter> | Sets or displays the device ID to value l, from A to Z (RS-485 only) | ID<enter> | Displays the current device ID setting | *ID*<enter> | 1, 2, 3 | Resets the ID to None | *ID*<enter> | 1, 2, 3 | Resets the ID to None
ID<enter> | Displays the current device ID setting | ID(n)<enter> | Sets or displays the device ID to value n, from 0 to 9 | IDn<enter> | Sets the device ID to value n, from 0 to 9 NOTE: When the ID feature is enabled, all commands to the device must be prefaced by the ID. Entering ID* clears the ID. | ID/<enter> | 1, 2, 3 | Sets the ID of the actuator to / (Must be 0-9 or A-Z) | ID/<enter> | 1, 2, 3 | Sets the ID of the actuator to / (Must be 0-9 or A-Z)
IDn<enter> | Sets the device ID to value n, from 0 to 9 NOTE: When the ID feature is enabled, all commands to the device must be prefaced by the ID number. Entering ID* disables this feature (discussed below.)
[6 empty] | IFMn<enter> | 1, 2, 3 | Sets the interface response mode to [0] no response string when an action command is sent, [1] basic response to action commands, or [2] extended response to action commands (needed for BCD interface) | IFMn<enter> | 1, 2, 3 | Sets the interface response mode to [0] no response string when an action command is sent, [1] basic response to action commands, or [2] extended response to action commands (needed for BCD interface)
IN<enter> | Starts a re-inialization sequence | IN<enter> | Starts an initialization sequence
[6 empty] | LGn<enter> | 1, 2, 3 | Sets the interface response mode to [0] no extra characters, or [1] longer response string to be fully backwards compatible with VICI microelectric actuators | LGn<enter> | 1, 2, 3 | Sets the interface response mode to [0] no extra characters, or [1] longer response string to be fully backwards compatible with VICI microelectric actuators
[6 empty] | LRN<enter> | 1 | Forces the actuator to find the stops on a newly installed valve. (Note: The valve must be installed before this command is sent.) | LRN<enter> | 1 | Forces the actuator to find the stops on a newly installed or reinstalled valve. (Note: A valve must be installed before this command is sent.)
|  |  |  | NP<enter> | Displays the number of positions the actuator is currently set to index | NP<enter> | 1, 2, 3 | Displays the number of positions the actuator is currently set to index | NP<enter> | 1, 2, 3 | Displays the number of positions the actuator is currently set to index
|  |  |  | NPnn<enter> | Sets the number of positions (nn) for the current valve | NPnn<enter> | 2 | Sets the number of ports (nn) for the current valve. (Must be an even number between 2 and 40) | NPnn<enter> | 2 | Sets the number of ports (nn) for the current valve. (Must be an even number between 2 and 40)
[6 empty] | NPnn<enter> | 3 | Sets the number of positions (nn) for the current valve. (Must be an even number between 2 and 40) | NPnn<enter> | 3 | Sets the number of positions (nn) for the current valve. (Must be an even number between 2 and 40)
SB<enter> | Displays the current baud rate | SB(n)<enter> | Sets or displays the baud rate, where n = 2400 to 115,220 (Default is 9600) | SB<enter> | Displays the current baud rate | SB<enter> | 1, 2, 3 | Displays the current baud rate | SB<enter> | 1, 2, 3 | Displays the current baud rate
SBnnnn<enter> | Sets the baud rate to 1200, 2400, 4800, 9600 (default), 14400, 19200, 28800, or 38400. The parity setting, number of data bits, and number of stop bits cannot be changed. (See section entitled Setting a New Baud Rate on the next page) |  |  | SBnnnn<enter> | Sets the baud rate to 1200, 2400, 4800, 9600 (default), 14400, 19200, 28800, or 38400. The parity setting, number of data bits, and number of stop bits cannot be changed. (See section entitled Setting a New Baud Rate on the next page) | SBnnnn<enter> | 1, 2, 3 | Sets the baud rate to 48(00), 96(00), 192(00), 384(00), 576(00), or 1152(00). The parity setting, number of data bits, and number of stop bits cannot be changed. | SBnnnn<enter> | 1, 2, 3 | Sets the baud rate to 48(00), 96(00), 192(00), 384(00), 576(00), or 1152(00). The parity setting, number of data bits, and number of stop bits cannot be changed.
|  |  |  | SD<enter> | Displays the digital input status where 0 = enabled and 1 = disabled
|  |  |  | SDn<enter> | Sets the digital input status to [0] BCD, [1] disabled, [2] parallel, or [3] binary. NOTE: Setting the status to [1] locks out communication through the manual controller and remote control cables, so digital input status defaults to [0] at power up.
|  |  |  | SL<enter> | Displays current Data Latch signal status
|  |  |  | SLn<enter> | This command displays or changes the requirement for a Data Latch signal to accompany BCD inputs. When set to [0] (factory default), the data latch is required for BCD inputs. When set to [1], the data latch is NOT required. This feature can reduce the number of control lines required for a system with a dedicated BCD output port and only one actuator connected. NOTE: Be sure all the BCD inputs are asserted within 20 milliseconds of each other or the actuator may be misdirected.
SM<enter> | Displays the current digital input mode (See section entitled Digital Control of the Actuator, subhead Input Modes) | SM(n)<enter> | Sets or displays the digital input mode, where n = 1 to 4 | SM<enter> | Displays the current default rotational direction for the digital inputs. | SM<enter> | 3 | Displays the current default rotational direction | SM<enter> | 3 | Displays the current default rotational direction
SMn<enter> | Sets the digital input mode to Mode n, where n is 1 or 2 |  |  | SMl<enter> | Sets the default rotational direction for the digital inputs to [F] for forward rotation, [R] for reverse rotation, or [A] to automatically choose the shortest route. This feature can be used instead of external wiring tied to the Manual and/or Autodirection input lines on the Digital I/O port. | SMl<enter> | 3 | Sets the default rotational direction to [F] for forward rotation, [R] for reverse rotation, or [A] to automatically choose the shortest route. | SMl<enter> | 3 | Sets the default rotational direction to [F] for forward rotation, [R] for reverse rotation, or [A] to automatically choose the shortest route.
SOnnnn<enter> | Turns off the position outputs after a delay, set in milliseconds to the closest 5 ms interval, from 0 to 30,000 ms. The outputs are always on (SO=0) by default. | SO(n)<enter> | Sets or displays the delay time before the position outputs are turned off, set in milliseconds to the closest 5 ms interval, from 0 to 30,000 ms. The outputs are always on (SO=0) by default. | SO<enter> | Displays the current offset value | SO<enter> | 1, 2, 3 | Displays the current offset value | SO<enter> | 3 | Displays the current offset value
|  |  |  | SOnn<enter> | Sets the offset value of the first position to be any number from 1 to 96 minus the total number of positions. Example: for a 10 position valve, the offset can be set from 1 to 86. (See section entitled Using the Offset Feature beginning on the next page.) | SOnn<enter> | 1, 2, 3 | Sets the offset value of the first position to be any number from 1 to 96 minus the total number of positions. Example: for a 10 position valve, the offset can be set from 1 to 86. | SOnn<enter> | 3 | Sets the offset value of the first position to be any number from 1 to 96 minus the total number of positions. Example: for a 10 position valve, the offset can be set from 1 to 86.
[6 empty] | STAT<enter> | 1, 2, 3 | Displays the status of the actuator | STAT<enter> | 1, 2, 3 | Displays the status of the actuator
[6 empty] | TM <enter> | 1, 2, 3 | Displays the amount of time required for the previous move, in milliseconds | TM <enter> | 1, 2, 3 | Displays the amount of time required for the previous move, in milliseconds
TO<enter> | Toggles the actuator to the opposite position | TO<enter> | Toggles the actuator to the opposite position |  |  | TO<enter> | 1, 2 | Toggles the actuator to the opposite position | TO<enter> | 1, 2 | Toggles the actuator to the opposite position
TT<enter> | Toggles the actuator to the opposite position, waits a preset delay time, then rotates back to the original position. | TT<enter> | Toggles the actuator to the opposite position, waits a preset delay time, then rotates back to the original position. |  |  | TT<enter> | 1, 2 | Toggles the actuator to the opposite position, waits a preset delay time, then rotates back to the original position. | TT<enter> | 1, 2 | Toggles the actuator to the opposite position, waits a preset delay time, then rotates back to the original position.
VR<enter> | Displays the part number and date of the firmware | VR<enter> | Displays the part number and date of the firmware | VR<enter> | Displays the part number and date of the firmware | VRn<enter> | 1, 2, 3 | Displays the current firmware version for [NULL], the main PCB, or [2] the serial interface PCB. | VRn<enter> | 1, 2, 3 | Displays the current firmware version for [NULL], the main PCB, or [2] the serial interface PCB.